## Slide 1: FANTOM II: Six-Month Clinical and Angiographic Results With a Radiopaque Desaminotyrosine Polycarbonate-Based Sirolimus-Eluting Bioresorbable Vascular Scaffold in Patients With Coronary Artery Disease
Dr. Alexandre Abizaid
Instituto Dante Pazzanese de Cardiologia
Sao Paulo, Brazil
On behalf of the FANTOM II investigators
FANTOM II

## Slide 2: Disclosure Statement of Financial Interest
Research Grants
Consulting Fees/Honoraria
Abbott Vascular
Boston Scientific
Elixir Medical
Medtronic
REVA Medical
Within the past 12 months, I or my spouse/partner have had a financial interest/arrangement or affiliation with the organization(s) listed below.
Affiliation/Financial Relationship
Company

## Slide 3: Fantom Bioresorbable Scaffold
Fantom® (REVA Medical) Sirolimus-Eluting Bioresorbable Scaffold
Desaminotyrosine Polycarbonate
  Key Scaffold Features
  Complete scaffold visibility under x-ray
  Single-step continuous inflation
  Clinically significant expansion range
  Good radial strength at 125 µm thickness
  Vasomotion restoration ~1 year (Preclinical)
  No special storage or handling
Deliverability
Vessel Patency
Visibility

## Slide 4: FANTOM II Study Investigators
Australia
  Dr. Muller, Dr. Jepson, Dr. Walters
Belgium
  Dr. De Bruyne
Brazil
  Dr. Abizaid, Dr. Costa, Dr. Chamie, Dr. Perin
Denmark
  Dr. Christiansen, Dr. Lassen, Dr. Okkels-Jensen
France
  Dr. Carrié, Dr. Chevalier, Dr. Fajadet, Dr. Collet
Germany
  Dr. Weber-Albers, Dr. Naber, Dr. Achenbach, Dr. Frey, Dr. Lutz, Dr. Kische, Dr. Ince, Dr. Brachman
Netherlands
  Dr. Amoroso, Dr. Wykrzykowska, Dr. Daemen
Poland
  Dr. Dudek, Dr. Kochman, Dr. Koltowski, Dr. Lesiak, Dr. Wojdyla

## Slide 5: FANTOM II Study Design and Endpoints
Study Design
  Safety and Performance Trial
  2.5mm to 3.5mm vessels
  Lesion length ≤ 20mm
  Primary Endpoint
    MACE & Late Loss at 6 Months
  Secondary Endpoints
    MACE all time points
    Late Loss at 9 Months
    Serial imaging sub-studies
      Cohort A: 24 months
      Cohort B: 48 months
Study Population
N= 240 Patients
28 Clinical Centers Participating
Cohort A
(117 Patients)
6 Mo Clinical
Follow-up (MACE)
Cohort B
(123 Patients)
6 Mo Clinical
Follow-up (MACE)
6 Mo Angio
(Late Loss)
OCT Sub-Study
9 Mo Angio
(Late Loss)
Includes OCT & IVUS
Sub-study @ 48 months
Annual Clinical
Follow-up (5 yrs)
Annual Clinical
Follow-up (5 yrs)

## Slide 6: FANTOM II – Cohort A & B Study Overview and Baseline Characteristics
Patient Characteristics (N=240) | 
Patient Age (mean ± SD) | 62.7 ± 10.1
Male | 70.4%
Diabetes | 23.8%
Current/Former Smoker | 59.6%
Hypertension | 73.3%
Hyperlipidemia | 70.4%
Prior PCI | 43.8%
Prior CABG | 2.9%
Prior MI | 26.3%
Recent LVEF <40% | 0.4% (N=231)
Study Population
N= 240 Patients
28 Clinical Centers
6 & 9 Month Follow-up
Imaging
6 Month Follow-up
Clinical
Angiographic
(cohort A N=100)
(cohort B ongoing)
OCT
(cohort A N=73)
(cohort B ongoing)
IVUS
(cohort A N=45)
(cohort B ongoing)
Long Term Follow-up
Clinical
(annual through 5 years)

## Slide 7: FANTOM II – Cohort A & B Lesion Characteristics and Procedural Outcomes
Target Lesion Location (n=240) | 
LAD | 48.8% (117)
LCX | 30.8% (74)
RCA | 20.4% (49)
ACC/AHA Lesion Class (n=238)* | 
Type A | 19.3% (46)
Type B1 | 48.7% (116)
Type B2 | 29.4% (70)
Type C | 2.5% (6)
Lesion Characteristics
Initial Outcomes
Acute Procedural Outcomes | 
Acute Technical Success (1) | 95.8%
Acute Procedural Success (2) | 99.1%
Clinical Procedural Success (3) | 99.6%
Defined as successful delivery and deployment of the intended scaffold in the intended lesion without device related complications.
Defined as acute technical success (see definition above), resulting in a residual stenosis of ≤50 percent with no immediate (in-hospital) MACE.
Defined as acute procedural success (see definition above), with no MACE thirty days post-intervention and with a final diameter stenosis ≤50 percent.
*As assessed by an independent core lab

## Slide 8: FANTOM II – Cohort A & B Safety Results
6 Month MACE Results Timeframe | Event N, (%)
In-Hospital | 2 (0.8%) (Post Procedure MI)
30-Day Follow-up | 1 (0.4%) (MI/TLR/SAT)
90-Day Follow-up | 0 (0%)
6-Month Follow-up | 2 (0.8%) (1 cardiac death, 1 TLR)
Components of the 6-Month Primary Endpoint (modified ITT): non-Hierarchical | N=240 N, (%)
MACE* | 5 (2.1%)
Cardiac Death | 1 (0.4%)
Target vessel MI | 3 (1.3%)
Clinically Driven TLR | 2 (0.8%)
* As adjudicated by an independent Clinical Events Committee
Note: Excludes 3 Non-Clinically Driven TLRs

## Slide 9: FANTOM II – Cohort A Angiographic – QCA Results*
In-Scaffold Analysis | Baseline (n=115) | Post Procedure (n=112) | 6 Months (n=100)
RVD (mm) | 2.68 ± 0.37 | 2.75 ± 0.40 | 2.70 ± 0.36
MLD (mm) | 0.79 ± 0.29 | 2.47 ± 0.37 | 2.23 ± 0.41
Diameter Stenosis (%) | 70.3 ± 10.4 | 7.6 ± 9.7 | 15.3 ± 15.2
Acute Gain (mm) |  | 1.67 ± 0.41 | 
Acute Recoil (%) |  | 2.9 ± 8.8 | 
Mean LLL (mm) |  |  | 0.5 ± 0.38
Median LLL (mm) |  |  | 0.2 (-0.43, 2.62)
In-Segment Analysis |  |  | 
Mean LLL (mm) |  |  | 0.17 ± 0.34
Median LLL (mm) |  |  | 0.15 (-0.43, 2.47)
Notes: Binary restenosis at 6 months follow-up = 2.0%
Data analyzed by an independent QCA core lab (Yale Cardiovascular Research Group, New Haven, US)
0.25 + 0.40

## Slide 10
Scaffold Lumen Area
 | Baseline | Follow-up | Difference | p-value
Mean lumen area (mm2) | 6.8 (1.7) | 5.7 (1.4) | -1.1 (-1.3;-0.9) | <0.0001
Minimal lumen area (mm2) | 5.3 (1.4) | 4.4 (1.4) | -1.0 (-1.3;-0.7) | <0.0001
Mean luminal area, mm2
Baseline
Follow-up
Mean(SD)
Reference: As presented by Jo Simonsen TCT 2016
PCI Research
Aarhus University Hospital
Notes: Descr. pics, + graph : Matched patients are interconnected
Expected due to healing response around the stented segment. This decrease in lumen area was comparable to what has been observed for both the ABSORB and the DESolve stent.

## Slide 11
Stent Area
 | Baseline | Follow-up | Difference | p-value
Mean stent area (mm2) | 7.1 (1.5) | 7.2 (1.4) | 0.1 (-0.02;0.24) | 0.12
Minimal stent area (mm2) | 5.9 (1.3) | 6.0 (1.3) | 0.1 (-0.02;0.25) | 0.08
Mean stent area, mm2
Baseline
Follow-up
Mean(SD)
Reference: As presented by Jo Simonsen TCT 2016
PCI Research
Aarhus University Hospital
Notes: Stent area was assessed based on the abluminal stent contour. We observed no evidence of stent area shrinkage … .
The reduction in molecular weitgh we start to see at the 6 months time point did importantly not affect the stent area.

## Slide 12
Strut Coverage
 | Follow-up
Covered struts | 98.1% (95.9;99.4)
Median(IQR)
Reference: As presented by Jo Simonsen TCT 2016
PCI Research
Aarhus University Hospital
Notes: Strut coverage was assessed on strut level and estimated as percentage of covered struts per recording. We observed a high coverage of stent struts after 6 months.
A standard limitation of OCT it is not possible to confidently detect layers of coverage below 10-15 um. Still 98% covereage is a good result. An even better coverage cannot be excluded. The meassure hwew is OCT coverage.

## Slide 13: FANTOM Program Clinical Summary
Fantom offers new and clinically important features
  Radiopacity
  Deliverability
  Single-step inflation
  No special handling requirements
Initial clinical data demonstrates
  Good acute performance
    Enhanced device deliverability
    Minimal residual stenosis and acute recoil (3%)
  Sustained performance and safety through 6 months
      Low MACE Rate (2.1%)
      Low late lumen loss (0.25mm)