## Slide 1
Horst Sievert
F Mahfoud, J Renkin, S Bertog, S Ewen, M Böhm, J-P Lengele, Y Honda, D Fleischmann, M Jaff, I Weinberg, M Hochul, W Wojakowski, A Schmid, RE Schmieder, M van der Giet, M de Bucourt, A Zirlik, P Neuzil, N Haratani, H Parise, T Fischell, A Persu
Chemical Renal Denervation with Alcohol:
The Post-Market Study Data and
The Target BP Program
CRT 2019
Washington DC, USA, March 4, 2019
.
10min
CardioVascular Center Frankfurt - CVC, Frankfurt, Germany

## Slide 2
Physician name | Company | Relationship
Horst Sievert | 4tech Cardio, Abbott, Ablative Solutions, Ancora Heart, Bavaria Medizin Technologie GmbH, Bioventrix, Boston Scientific, Carag, Cardiac Dimensions, Celonova, Comed B.V., Contego, CVRx, Dinova, Edwards, Endologix, Hemoteq, Lifetech, Maquet Getinge Group, Medtronic, Mitralign, Nuomao Medtech, Mokita, Occlutech, pfm Medical, Recor, Renal Guard, Rox Medical, Terumo, Vascular Dynamics, Venus, Vivasure Medical | Consulting fees, Travel expenses, Study honoraria to institution
Disclosures

## Slide 3: Participating Clinical Centers
Site Name City, Country | Principal Investigator
Clinique Universitaires Saint‑Luc Brussels, Belgium | Alexandre Persu
Oddzial Kardiologii Inwazyjnej Tychy, Poland | Marius Hochul
CardioVascular Center CVC Frankfurt, Germany | Horst Sievert
Klinik für Innere Medizin III Homburg/Saar, Germany | Felix Mahfoud
Universitätsklinik Erlangen Erlangen, Germany | Roland Schmieder
Charité‑Universitätsmedizin Berlin Berlin, Germany | Marcos Bucourt
Universitäts‑Herzzentrum Freiburg, Germany | Andreas Zirlik
Na Homolce Hospital Prague, Czech Republic | Petr Neužil
Elisabeth‑Krankenhaus Essen GmbH Essen, Germany | Thomas Schmitz

## Slide 4: Committees and Core/Central Laboratories
Data Safety Monitoring Board
  Cardiovascular Research Foundation: Bernard Gersh, MD; Harold Lee Dauerman, MD; Phyllis August, MD.
Clinical Event Committee
  Cardiovascular Research Foundation: Steven Marx, MD; Clive Rosendorff, MD; Michele Mokrzycki, MD
Ambulatory Blood Pressure Core Lab:
  ERT (Formarly Biomedical Systems)
MRA/CTA/Angio Core Lab
  Cardiovascular Core Analysis Laboratory
Renal Duplex Ultrasound
  Vascore
Drug Compliance (Urine) Central Lab
  Synlab

## Slide 5: Alcohol-Mediated Denervation using the Peregrine system
Micro-volume (0.3 mL−0.6 mL) infused directly to the perivascular region
Extracellular fluid helps spread alcohol circumferentially in the perivascular region
Alcohol activity range self-limited through dilution by extracellular fluid
Perivascular Sites Where Device Infuses Alcohol
Expanded View of Device
Infusing Alcohol
CAUTION: The Peregrine System kit is currently being studied to evaluate safety and efficacy for renal denervation (also referred to as chemical renal denervation) It is an investigational product not approved in the US. Its use is limited to investigational use in clinical trials. The Peregrine System Infusion Catheter is CE marked, and only for use in the European Union with a neurolytic agent (i.e. alcohol) for treatment of systemic HTN.
PVRD Chapter 13; RENAL DENERVATION: A New Approach Treatment – Page 107-116;2015

## Slide 6: Patient Treatment Using the PeregrineTM Catheter (0.6 mL Alcohol)
CAUTION: The Peregrine System kit is currently being studied to evaluate safety and efficacy for renal denervation (also referred to as chemical renal denervation) It is an investigational product not approved in the US. Its use is limited to investigational use in clinical trials.
The Peregrine System Infusion Catheter is CE marked, and only for use in the European Union with a neurolytic agent (i.e. alcohol) for treatment of patients with systemic hypertension.

## Slide 7: Study Objective And Study Primary Endpoints
Study objective: collect additional safety and performance data using the Peregrine System™ and a neurolytic agent (i.e. alcohol)
Primary efficacy endpoint: reduction of the 24-hour mean systolic ABPM compared to baseline, assessed at 6-month follow-up
Primary safety endpoint: Major Adverse Event (MAE) at 1 month
  Peri-procedural major vascular complications
  Major bleeding as defined by the TIMI Bleeding Classification
  Acute kidney injury (AKI)
  Death of any cause
Notes: Acute kidney injury (Modified RIFLE Classification):
Change in Serum Creatinine (up to 72 hours) compared to baseline
  Stage 1: Increase in serum creatinine to 150-200% (1.5-2.0 times increase compared with baseline) or increase of ≥0.3 mg/dl (≥26.4 μmol/L)
  Stage 2: Increase in serum creatinine to 200-300% (2.0-3.0 times increase compared with baseline)
  Stage 3*: Increase in serum creatinine to ≥ 300% (>3 times increase compared with baseline) or serum creatinine of ≥4.0 mg/d (≥354 μmol/L) with an acute increase of at least 0.5 mg/dl (44 μmol/L)
  Subjects receiving renal replacement therapy are considered to meet Stage 3 criteria irrespective of other criteria.
Major Complications
  Vascular repair / need for vascular repair at the femoral artery access site
  Bleeding requiring transfusion
  New permanent ipsilateral lower extremity ischemia
  Permanent nerve injury at access site
  Pseudo-aneurysm requiring treatment with ultrasound-guided compression or fibrin/thrombin injection
  Access site-related hematoma greater than 6 cm requiring vascular repair or intervention other than manual compression including ultrasound guided compression or ultrasound guided injection of prothrombotic materials
  Major Artery Dissection (see Artery Dissections’ Definition)
  Minor dissections are dissections that do not have any flow-limitation of blood supply to an organ as assessed by blood tests and/or clinical observations, and where further intervention is not performed. These types of dissections are considered benign and usually heals over 3 to 6 months without sequelae.

## Slide 8: Key Inclusion Criteria
Age 18-80
Renal artery diameter 4-7mm, and length ≥ 5mm
Patient has hypertension with:
  triplicate of office systolic blood pressure (BP) with a mean of ≥150 mmHg AND office diastolic BP of ≥85 mmHg
  a 24-hour mean systolic ABPM ≥135 mmHg with ≥70% valid readings
  is receiving and adhering to a stable medication regimen of at least three anti-hypertensive medications of different classes (for at least 4 consecutive weeks), one of which must be a diuretic

## Slide 9: Key Exclusion Criteria
Contraindication known for conventional percutaneous interventional procedures;
Documented severe untreated obstructive sleep apnea (AHI ≥30 per hour);
Type 1 diabetes mellitus;
eGFR ≤20 mL/min/1.73m2, based on the CKD-EPI equation
Imaging-assessed renal artery anatomy abnormalities or variations based on Investigator’s evaluation of the screening images
Patient is known to have a non-functioning kidney or unequal renal size (>2 cm difference in renal length between kidneys associated with a chronic kidney disease or a deterioration of the kidney function);
History of myocardial infarction, unstable angina pectoris, or stroke/TIA within the last six months from planned procedure;
Heart failure (NYHA III or IV) or has an ejection fraction ≤30%;
Chronic atrial fibrillation

## Slide 10: Patient Disposition
Enrollment / Screening N=114
Procedure N=45
1 Month N=45
3 Months N=44
6 Months N=44
Patient Withdrew Consent N=1
Screen Failure N=59
Patient Withdrew Consent N=8
Other / Missing Reason N=2
12 Months N=29
Visit Not Yet Completed/Due N=13
Patient Withdrew Consent N=1
Other N=1

## Slide 11: Baseline Characteristics (n = 45)
Characteristic | 
Age, years (mean ± SD) | 55 ± 10
Male | 62%
Chronic Kidney Disease | 2%
Diabetes (Type II) | 33%
Arrhythmia | 13%
Hyperlipidemia | 49%
Coronary Artery Disease | 9%
Chronic Obstructive Pulmonary Disease | 9%
BMI, kg/m2 (mean ± SD) | 30.7 ± 5.8
Serum Creatinine, mg/dl (average ± SD) | 0.92 ± 0.19
eGFR, per CKD-EPI, mL/min/1.73m2 (mean ± SD) | 85 ± 16
Data as of 11 Dec 2019
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
Notes: 6 arrythmia: 3 AF, 1 RBB/Sinus Bradycardia, 1 SVES (Supraventicular Extra-systolie), and 1 supraventricular tachycardia

## Slide 12: Key Procedural Safety and Performance (n = 45)
Parameter | 
Procedural device deficiencies | 0%
Procedure completed as intended | 100%
Peri-procedural serious adverse events* | 
Vascular access site pseudoaneurysm** | 2%
Vascular access site Hemorrhage** | 2%
Vascular access site infection | 2%
Abdominal pain | 2%
Renal artery complication requiring intervention*** | 0%
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
* As adjudicated by the CEC to be related to procedure, if available, otherwise by investigator
** Pseudoaneurysm requiring intervention and bleeding reported in one patient.
***Two minor dissections not requiring intervention were reported. Both were resolved without sequelae.
Data as of 11 Dec 2019
Notes: Two minor dissections:
One was due to user error: the operator inadvertently moved the catheter forward before retracting the needles. The dissection did not require any intervention. Renal function was not changed per-procedural and during follow-up. CTA at 6 months done and analyzed by the core lab who found no change from the baseline (pre-procedure) images, confirming the complete resolution of the dissection.
One was identified post-hoc during the CL review and analysis of the procedural images taken immediately following the alcohol infusion. The images showed a very short intimal flap (<1mm) minimally protruding into the vessel lumen at the location of the preceding needle penetration. This event is considered not clinically significant by the operator since there is no indication of a progressive distal dissection or of a luminal narrowing. Diagnostic Renal Duplex Ultrasound at 6 months showed no renal artery stenosis confirming complete resolutions of this minor dissection.

## Slide 13: Medication through 6 Months Patient Reported
 | Baseline | 1 Month | 3 Month | 6 Month
Mean Antihypertensive Medications |  |  |  | 
N | 45 | 45 | 44 | 44
Mean ± SD | 5.1 ± 1.5 | 4.8 ± 1.5 | 4.8 ± 1.4 | 4.7 ± 1.6
Number of Antihypertensive Medications |  |  |  | 
3 | 24% | 22% | 23% | 23%
4 | 11% | 16% | 16% | 11%
5 | 20% | 22% | 25% | 25%
6 | 22% | 27% | 23% | 23%
7 | 22% | 11% | 11% | 11%
Number of Antihypertensive Medications |  |  |  | 
Increased 1 | - | 2% | 2% | 5%
No Change | - | 84% | 84% | 73%
Decreased 1 | - | 2% | 5% | 9%
Decreased 2 | - | 7% | 5% | 9%
Decreased 3 | - | 4% | 5% | 5%
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
Data as of 11 Dec 2019

## Slide 14: 1, 3, 6 and 12 Months Systolic Blood Pressure Reduction*
Responders**: 69% 58% 61% 68% 64% 67% 71% 54%
* 24-hour data from CL; OBP Data: site reported. 12 Month data not yet monitored
** Responders are defined as ≥ 10mmHg drop for OBP and ≥ 5mmHg drop for ABPM
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
Data as of 11 Dec 2019
Chart: 
| 1-M (N=45) | 3-M (N=43) | 6-M (N=44) | 12-M (N=28)
 | -18.0222 | -17.5814 | -18.2955 | -20.3929
Chart: 
| 1-M (N=42) | 3-M (N=36) | 6-M (N=42) | 12-M (N=26)
 | -8.5476 | -9.7222 | -11.1429 | -10.4231

## Slide 15: 1, 3, 6 and 12 Months Diastolic Blood Pressure Reduction*
* 24-hour data from CL; OBP Data: site reported. 12 Month data not yet monitored
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
Data as of 11 Dec 2019
Chart: 
| 1-M (N=42) | 3-M (N=36) | 6-M (N=42) | 12-M (N=26)
 | -5.5476 | -6.4167 | -6.5 | -7.2308
Chart: 
| 1-M (N=45) | 3-M (N=43) | 6-M (N=44) | 12-M (N=28)
 | -9.2444 | -8.4651 | -9.6591 | -10.8571

## Slide 16: 6-Month Individual Patients Changes in BP
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union
Data as of 11 Dec 2019
Chart: Systolic OBP
 | -92 | -52 | -45 | -42 | -40 | -40 | -40 | -38 | -35 | -35 | -34 | -34 | -29 | -29 | -29 | -25 | -24 | -20 | -19 | -18 | -18 | -18 | -16 | -15 | -14 | -14 | -11 | -8 | -7 | -5 | -5 | -4 | -3 | -3 | -2 | 4 | 4 | 5 | 5 | 6 | 6 | 6 | 10 | 12 | 0
Chart: 24-Hr Systolic ABPM
 | -63 | -42 | -34 | -23 | -23 | -21 | -21 | -21 | -20 | -20 | -18 | -18 | -14 | -14 | -13 | -12 | -12 | -11 | -11 | -11 | -11 | -10 | -9 | -8 | -8 | -8 | -7 | -7 | -6 | -6 | -3 | -3 | -2 | -1 | 0 | 0 | 3 | 3 | 5 | 9 | 10 | 13 | 0 | 0 | 0

## Slide 17: The TARGET BP Program
 | TARGET BP OFF MED | TARGET BP I (2-5 Meds) | TARGET BP Safety “Safety Data”
N | 90 | 300 | 300 (including crossover)
Number of Sites/Geography | EU only | US and OUS | US and OUS
Study Design | RCT, blinded, 1:1 (Tx vs. sham) POC (phase II) | RCT, blinded, 1:1 (Tx vs sham), phase 3 pivotal study | Open label, prospective, study (including the cross over subjects)
Primary Endpoint - Efficacy | Mean Reduction in SPB by ABPM at 8 weeks | Mean Reduction in SPB by ABPM at 3 Months | Mean Reduction in SPB by ABPM at 3 Months
Secondary Major Safety Endpoint | MAE at 1 month | MAE up to 6 M | MAE up to 6 M
Unblinding | Post-6 M visit by all pts | Post-6 M visit by all pts | NA
Crossover | After unblinding, and review by DSMB | After unblinding, 1-year data reviewed by DSMB | NA
Follow-up Duration | 2 years | 3 years | 1 year

## Slide 18: Conclusions
Alcohol-mediated renal denervation allows efficient and safe nerve ablation and results in significant BP lowering
TARGET BP OFF-MED (Europe) phase II study will provide further validation of this approach (enrolling)
The randomized on-med controlled trial (TARGET BP I) is cleared by the FDA
The Peregrine System™ Infusion Catheter is CE marked and only for use in the European Union

## Slide 19: Thank you!
HorstSievertMD@aol.com