## Slide 1
FFRangio Accuracy vs. STandard FFR:
Results from the FAST-FFR Trial
William F. Fearon, MD
On Behalf of the FAST-FFR Study Investigators

## Slide 2
Disclosure Statement of Financial Interest
Affiliation/Financial Relationship Institutional Grant/Research Support: Consulting Fees/Honoraria:
Major Stock Shareholder/Equity Interest: Royalty Income:
Ownership/Founder: Salary:
Intellectual Property Rights: Other Financial Benefit:
Company
Abbott, Medtronic, CathWorks, Boston Scientific
Stock Options HeartFlow
Within the past 12 months, I or my spouse/partner have had a financial interest /arrangement or affiliation with the organization(s) listed below

## Slide 3: Fractional Flow Reserve
Distal Pressure (Pd)
Proximal Pressure (Pa)
FFR = Pd / Pa
during maximal flow
Pd
Pa
Pd / Pa = 60 / 100 FFR = 0.60

## Slide 4: FFRangio
Optimal projections Optimal frame Motion compensation
Optimal 2D angiography
1
Extracting centerlines Tree topology
3D model reconstruction
2
Bifurcation analysis 2D-QCA analysis Estimating diameters
Stenosis assessment
3
Resistance mapping Maximum blood flow Flow rate ratio
Hemodynamic evaluation
4

## Slide 5: Methods
Objective
The FAST-FFR study is a prospective, multicenter, international trial comparing the accuracy of on-site FFRangio with pressure wire-derived FFR.

## Slide 6: Methods
Co-Primary Endpoints
Sensitivity & Specificity of FFRangio as compared with pressure wire-derived FFR using a cutoff value ≤ 0.80.
Powered to meet the lower bound of the 95% CI for pre- defined performance goals set at:
  Sensitivity = 0.70
  Specificity = 0.75

## Slide 7
FFR and FFRangio Case Example
FFR=0.68

## Slide 8
FFR and FFRangio Case Example
FFR=0.87

## Slide 9: Results
Physiologic Result
FFR
Mean
0.81 ± (0.13)
Median
0.83 (0.74, 0.90)
% of positive lesions (≤ 0.80)
43.3%
% within 0.70-0.90
58.9%
% within 0.75-0.85
31.3%
FFRangio
0.80 ± (0.12)
0.82 (0.73, 0.89)
45.5%
63.6%
31.0%
FFR and FFRangio Results
FFRangio was successfully measured in 98.7% of cases

## Slide 10: Results Primary and Secondary Endpoints
Diagnostic Characteristic
Sensitivity 93.5% (87.8, 96.6)
Specificity 91.2% (86.0, 94.6)
Diagnostic accuracy 92.2% (88.7, 94.8)
Positive Predictive Value 89.0% (82.6, 93.2)
Negative Predictive Value 94.8% (90.3, 97.3)
Accuracy around FFR cutpoint (0.75-0.85)
Sensitivity 88.5%
Specificity 85.1%
Diagnostic accuracy 86.9%

## Slide 11: Summary
cases included.
FFR derived from routine coronary angiography (FFRangio) had very high sensitivity, specificity and diagnostic accuracy, all of which were greater than 90% for predicting the reference standard, coronary pressure wire-derived FFR.
FFRangio and FFR remained highly correlated over the entire range of FFR values.
FFRangio was successfully measured in almost all

## Slide 12: Conclusion
long-term outcomes.
FFRangio may provide an easier and potentially faster method for performing physiology guided assessment of the overall coronary angiogram with similar accuracy to the reference standard, coronary pressure wire-based FFR.
This may translate into a greater percentage of patients undergoing physiologic guidance for revascularization decisions and ultimately improve