## Slide 1: Coronary Tools and Tricks to Treat Paravalvular Leak
Nahid El Faquir, Anne-Marie Maugenest,
Mihai Strachinaru, Peter PT de Jaegere,
Nicolas M van Mieghem
Erasmus University Medical Center,
Rotterdam, The Netherlands

## Slide 2
Nahid El Faquir, MD
I have no relevant financial relationships

## Slide 3: Learning objectives
To illustrate coronary tools and tricks for paravalvular leak closure

## Slide 4: Clinical Presentation
57 year old male
Medical history:
  1981 Hodgkin’s Lymphoma
  09-2016 Severe aortic stenosis:
  SAVR Mechanical prosthesis 25mm
Current:
  NYHA III
  Holodiastolic murmur parasternal left ICS 3
  Normal kidney function, mild hemolysis and anemia (Hb 6,9 mmol/L)

## Slide 5: EKG

## Slide 6: TEE - PVL after SAVR

## Slide 7: CT – Periprosthetic PVL space

## Slide 8: Diagnostic 6F MP catheter & coronary wire

## Slide 9: Coronary balloon facilitated PVL crossing
Notes: Through diagnostic MP 6F

## Slide 10: Introducing pre-shaped Safari guide wire
Notes: To exchange the sheathless 7,5F JR4

## Slide 11: AV-Plug positioning through sheathless catheter
Notes: Plug positioning through sheathless 7,5F guide

## Slide 12: Final angio

## Slide 13: TEE - post AV-Plug implantation

## Slide 14: Take Home Messages
Coronary wires can simplify crossing paravalvular leak
Coronary balloons can provide additional support to help advance catheters through challenging (serpentine) PVL
Sheathless 7,5F guiding catheters can be excellent platforms to deliver plugs in PVL