## Slide 1: Clinical evaluation of the Hydra self-expanding THV: 30 days results from the GENESIS-II study
Nagendra Boopathy Senguttuvan, John Jose, Anmol Sonawane, Praveen Chandra
- on behalf of the GENESIS-II study Investigators

## Slide 2: Potential conflicts of interest
Speaker's name : Nagendra Boopathy Senguttuvan
☑ I have the following potential conflicts of interest to report:
☑ Proctor for Hydra valves

## Slide 3: Why this study?
The Hydra transcatheter aortic valve (THV) utilized in the GENESIS-II study is equipped with an active release mechanism for tentacle deployment, distinguishing it from the Hydra THV used in earlier studies.
In this regard, the primary objective of the GENESIS-II study is to assess the ongoing safety and performance of the Hydra THV in treating severe aortic stenosis in patients at high surgical risk.*
In addition, we also intended to compare the clinical outcomes of the GENESIS-II study with the historical control group of the GENESIS study^ where the Hydra THV was used.
*The high-risk classification was assigned to subjects with severe aortic stenosis symptoms for whom conventional aortic valve replacement surgery was associated with high-risk equivalent to an STS risk score that was ≥3%. Patients with an STS risk score that was <3% were assigned high risk if frailty indices and/or existing co-morbidities not captured by STS were also present.
^Catheter Cardiovasc Interv. 2021 Aug 1;98(2):371-379.

## Slide 4: Participating Sites
Study Principal Investigators:
Dr Praveen Chandra, Medanta–The Medicity, Gurugram, India
Dr John Jose,
Christian Medical College, Vellore, Tamil Nadu, India
Coordinating Investigators
Dr Nagendra Boopathy Senguttuvan Sri Ramachandra Hospital, Chennai, Tamil Nadu, India
Dr Anmol Sonawane
Breach Candy Hospital Trust, Mumbai, India
CTRI number:
CTRI/2021/06/034172
Sr. No. | Investigators | Sites (19 sites across India)
1 | Dr N B Senguttuvan | Sri Ramachandra Hospital, Chennai
2 | Dr Sandeep Bansal | Safdarjung Hospital, New Delhi
3 | Dr Rahul Gupta | Apollo Hospitals Enterprise Limited, Thane
4 | Dr John Jose | Christian Medical College, Vellore
5 | Dr Anoop Agrawal | Care Hospital, Hyderabad
6 | Dr C N Manjunath | Sri Jayadeva Institute, Bangalore, Karnataka
7 | Dr Sharad Chandra | KGM University, Lucknow
8 | Dr Suhas Hardas | Sahyadri Super Speciality Hospital, Pune
9 | Dr Milan Chag | CIMS Hospital Pvt. Ltd., Ahmadabad
10 | Dr R K Premchand | KIMS Ltd., Hyderabad
11 | Dr Shafeeq Mattummal | Aster MIMS, Kozhikode
12 | Dr Rajpal Abhaichand | GKNM Hospital, Coimbatore
13 | Dr Bishav Mohan | D M College and Hospital, Ludhiana
14 | Dr G Sengottuvelu | Apollo Hospitals, Chennai
15 | Dr Ranjan Shetty | Manipal Hospitals, Bangalore
16 | Dr Saibal Mukhopadhyay | G B Pant Hospital, New Delhi
17 | Dr Tejas Patel | Apex Heart Institute, Ahmadabad
18 | Dr Niteen Deshpande | Spandan Heart Institute, Nagpur
19 | Dr Abhilash S P | Sree Chitra, Thiruvananthapuram

## Slide 5: What did we study?
Three Tentacle Design More flexibility to navigate through aortic arch
Extended Internal Sealing Skirt
Mitigates paravalvular leak
Non-Flare Inflow part Reduces conduction abnormalities
Supra-annular Valve
For optimized hemodynamics
Radio Opaque Frame Markers
Accurate positioning
HYDRA THV - Self-expanding, Supra-annular, Bovine pericardium
3 sizes studied
22 mm 26 mm 30 mm
Large Frame Cells (≥15 F) Facilitates easy access to the coronary arteries

## Slide 6: What did we study?
Non-braided transparent capsule & flexible shaft
For better deliverability
14F inline sheath*
For smaller access
Improved control handle with active release feature
For ease of release
14F*
HYDRA New AVDC System – High Flexibility, Better Control, Up to 80% Recapturable
* 14F equivalent integrated sheath diameter for patients requiring 22 mm or 26 mm Hydra THV
14F* delivery system

## Slide 7: How was the study executed?
Study design:
Prospective, multicenter, non-randomized investigational study
Population:
Patients with severe aortic stenosis at high surgical risk
Primary safety endpoint:
Cardiovascular mortality at 30 days
Primary performance endpoint:
Acute device success at implantation*
Hemodynamic parameters analyzed:
Mean aortic gradient
Effective orifice area
*Device success defined as per Valve Academic Research Consortium (VARC)-2 criteria
6 months follow-up is ongoing
Between November 2021 and November 2023, a total of 40 patients were recruited from 19 sites across India
Patients screened (n=43)
Enrollment (n=40)
30-day follow-up (n=39)
Death (n=1)
Screening failure (n=3)

## Slide 8: How was the study executed?
Age ≥ 65 years
STS score ≥3%
STS score <3% with presence of frailty indices and/or existing co-morbidities
Aortic annulus diameter meets the range 17.0 mm-20.0 mm for Hydra 22, 20.0 mm-23.5 mm for Hydra 26 and 23.5 mm-27.0 mm for Hydra 30
Patient with or without >mild aortic regurgitation with echocardiography derived aortic valve area of <1.0 cm2 AND mean gradient >40 mmHg or peak velocity greater
>4.0 m/s or doppler velocity index (DVI) ≤0.25
Key inclusion criteria
Key exclusion criteria
Patients with congenital bicuspid aortic valve
Patients with associated severe mitral or tricuspid valvular regurgitation or moderate to severe mitral stenosis
Patient has aortic root diameter of < 26 mm or >36 mm
Patient has resting LVEF < 20%
Patients with significant coronary artery disease
Patient has echocardiographic evidence of intracardiac mass, thrombus, or vegetation
Cardiac shock (vasopressor dependence or mechanical support)
Patients with severe lung disease with FEV 1 < 50%
Patient with expected life expectancy less than 24 months
FEV: Forced expiratory volume, LVEF: Left ventricular ejection fraction

## Slide 9: What are the essential results?
Baseline Characteristics | N=40
Age, years | 74.4±6.7
Male, n (%) | 26 (65.0)
Body-mass index | 25.7±5
STS Score (%) | 3.5±1.6
Diabetes mellitus, n (%) | 17 (42.5)
Hypertension, n (%) | 24 (60.0)
Coronary artery disease, n (%) | 10 (25.0)
Previous CABG, n (%) | 5 (12.5)
Previous PCI, n (%) | 4 (10.0)
Peripheral artery disease, n (%) | 2 (5.0)
Chronic kidney disease, n (%) | 2 (5.0)
COPD, n (%) | 4 (10.0)
Prior stroke, n (%) | 1 (2.5)
Permanent pacemaker, n (%) | 2 (5.0)
NYHA III/IV functional class, n (%) | 18 (45.0)
Mean aortic gradient, mmHg | 53.1±12.5
Aortic valve area, cm2 | 0.6±0.2
≥ Moderate aortic regurgitation, n (%) | 8 (20.0)
Procedural Characteristics | N=40
Trans-femoral Access, n (%) | 40 (100.0)
Local anesthesia and conscious sedation, n (%) | 40 (100.0)
Implanted valve size Hydra 22 mm, n (%) Hydra 26 mm, n (%) Hydra 30 mm, n (%) | 2 (5.0) 22 (55.0) 16 (40.0)
Pre-dilatation, n (%) | 33 (82.5)
Post-dilatation, n (%) | 19 (47.5)
Procedure time, min | 116.8±51.3
Fluoroscopy time, min | 39.9±9.5
Contrast Volume, mL | 204.9±119.4
Duration of hospital stay, days | 4.0±2.4

## Slide 10: What are the essential results?
Primary safety endpoint
Cardiovascular mortality at 30 days
2.5%
Primary performance endpoint
Device success at implantation
95%*
*Device was successful in 38 out of 40 patients, with one subject experiencing peri-procedural death, while another required surgical re- intervention due to device embolization.
Clinical Outcomes (N=40) | In-hospital | 30-days
All-cause mortality, n (%) | 1 (2.5) | 1 (2.5)
Cardiovascular mortality, n (%) | 1 (2.5) | 1 (2.5)
Device-related mortality, n (%) | 1 (2.5) | 1 (2.5)
All stroke , n (%) | 0 (0) | 0 (0)
Disabling stroke, n (%) | 0 (0) | 0 (0)
Nondisabling/Transient ischemic attack, n (%) | 0 (0) | 0 (0)
Life-threatening bleeding, n (%) | 0 (0) | 0 (0)
Major bleeding, n (%) | 0 (0) | 0 (0)
Major vascular complication, n (%) | 0 (0) | 0 (0)
Acute kidney injury, n (%) | 0 (0) | 0 (0)
Myocardial infarction, n (%) | 0 (0) | 0 (0)
New permanent pacemaker implantation, n (%) | 3 (7.5) | 3 (7.5)
In pacemaker naïve patients, n (%) | 3 (7.9) | 3 (7.9)

## Slide 11: What are the essential results?
70
60
50
40
30
20
10
0
2,6
2,1
1,6
1,1
0,6
0,1
53.1±12.5
(49.1–57.1)
6.6±3.0
(5.5–7.7)
0.6±0.2
(0.5–0.7)
Hemodynamic Performance
Mean aortic gradient [mmHg]
EOA (cm2)
8.4±4.0
(7.1–9.7)
1.9±0.7
(1.7–2.1)
1.9±0.6
(1.7–2.1)
P<0.001 P<0.001
Pre procedure Post procedure 30 days
P<0.001
P<0.001
Mean aortic gradient [mmHg] EOA [cm2]
Progressive enhancement in effective orifice area has resulted in a significant improvement in the mean aortic valve gradient

## Slide 12: What are the essential results?
NYHA Functional Class (%)
65,8%
55,0%
35,0%
2,6%
31,6%
10,0%
Baseline 30 days
NYHA I NYHAII NYHAIII NYHA IV
Improvement in NYHA score by at least 1 and 2 functional class from baseline to 30 days occurred in 73.7% and 34.2% of patients, respectively.
Para-valvular Leak (%)
78,9%
89,5%
21,1%
10,5%
Post Procedure 30 days
None/Trace Mild Moderate Severe
No moderate or severe paravalvular leak was observed at 30 days follow-up
Six-minute walk test
Distance (m)
200
220
240
260
280
217.0 ± 50.9
275.9 ± 66.4
P<0.001
Pre-procedure 30 days
The six-minute walk test distance increased from 217.0±50.9 m at baseline to 275.9±66.4 m at 30 days (P<0.001)

## Slide 13: Why is this important?
Comparative clinical outcomes of the GENESIS-II study with the historical control group of the GENESIS study
Variables | GENESIS-II (n=40) | GENESIS (n=40) | p value
Device success as per VARC-2 | 95.0% | 92.5% | >0.999
All-cause mortality at 30 days | 2.5% | 10.0% | 0.178
Cardiovascular mortality at 30 days | 2.5% | 7.5% | 0.316
Major vascular complication at 30 days | 0% | 2.5% | 0.311
New permanent pacemaker implantation at 30 days | 7.5% | 7.5% | 0.982
Mean aortic-valve gradient at 30 days (mmHG) | 8.4±4.0 | 6.7±2.8 | 0.054
EOA at 30 days (cm2) | 1.9±0.6 | 2.3±0.6 | 0.004
Moderate/severe para-valvular leak at 30 days | 0% | 0% | -
6-minute walk test distance at 30 days, (m) | 275.9 ± 66.4 | 229±24 | <0.001

## Slide 14: The essentials to remember
The Hydra THV is the first device with an active release mechanism for deploying the supra-annular valve.
The observed high rate of device success and procedural safety in the GENESIS-II study confirm the safety and performance of the device in treating severe aortic stenosis in patients at high surgical risk.
The GENESIS-II study demonstrates higher device success rates, lower mortality, and absence of major vascular complications compared to a historical control group, while maintaining comparable hemodynamic outcomes.

## Slide 15: PCRonline.com