## Slide 1: Impact of hemodynamics on clinical outcomes after self-expanding transcatheter aortic valve implantation in failed surgical valves
David Holzhey, MD for the VIVA and EUS Investigators

## Slide 2: Potential conflicts of interest
Speaker's name : David Holzhey, MD
☑ I have the following potential conflicts of interest to declare:
Proctor and advisor for Medtronic, proctor for Boston Scientific and advisor for Edwards Lifesciences
David Holzhey, MD

## Slide 3: Background
Two trials, Valve in Valve (VIVA) in Europe and Expanded Use United States (EUS) have been combined to analyze TAV in SAV.
We evaluated clinical and hemodynamic outcomes of patients with failing surgical aortic valves (SAV) that underwent transcatheter aortic valve implantation (TAVI) with the self-expandable platform.
David Holzhey, MD

## Slide 4: Methods
VIVA and EUS were prospective, observational, single-arm, multi-center studies that evaluated the safety and effectiveness of TAVI in patients with symptomatic failed surgical aortic valves.1,2
This analysis combined data from these 2 studies to further assess outcomes of TAV in SAV procedures using a CoreValve or Evolut R bioprosthesis.
The primary safety endpoint was all-cause mortality analyzed by
  SAV failure mode (regurgitation, stenosis or combined)
  Post-implant mean gradient ≥20 mmHg
1Tchétché D, et al. JACC Cardiovasc Interv. 2019; 12, 923–32.
2Dauerman H, et al. Circ Cardiovasc Interv. 2019; 12(10): e008155
David Holzhey, MD

## Slide 5: Baseline and valve characteristics
Mean ± standard deviation or n (%) | N=425
Age, years | 78.2 ± 9.4
Male sex | 237 (55.8)
STS-PROM, % | 7.8 ± 6.1
Logistic EuroSCORE, % | 24.2 ± 15.5
Previous stroke | 39 (9.2)
Diabetes mellitus | 141 (33.2)
Previous MI | 66 (15.5)
Surgical valve age, years | 9.8 ± 4.4
Surgical valve type, stented | 372 (87.5)
Surgical valve label size (mm) | 
19-21 mm | 154/412 (37.4)
23 mm | 131/412 (31.8)
25–29 mm | 127/412 (30.8)
Surgical valve failure mode | 
Stenosis | 239 (56.2)
Regurgitation | 95 (22.4)
Combined | 91 (21.4)
David Holzhey, MD

## Slide 6: Clinical and haemodynamic outcomes
Kaplan-Meier rates at 2 years (%) | All | SAV Failure Mode |  | 
 |  | Regurgitation | Stenosis | Combined
All-cause mortality | 16.7 | 9.8 | 20.3 | 14.5
All stroke | 6.1 | 3.2 | 8.8 | 2.4
Disabling/major stroke | 2.1 | 1.1 | 3.4 | 0.0
Prosthetic valve thrombosis | 1.3 | 2.2 | 0.9 | 1.2
Prosthetic valve endocarditis | 0.8 | 1.2 | 1.0 | 0.0
Life-threatening bleeding | 7.5 | 7.5 | 7.6 | 6.8
Acute kidney injury stage II or III | 0.8 | 0.0 | 1.1 | 1.1
New pacemaker implantation | 12.5 | 11.9 | 14.0 | 9.1
Haemodynamics (mean ± SD) |  |  |  | 
Discharge mean gradient, mmHg | 18.5±9.7 | 14.5±6.9 | 20.1±10.3 | 18.3±9.2
Mean gradient at 1 year, mmHg | 15.5±8.0 | 12.8±8.1 | 16.8±7.6 | 15.7±8.1
David Holzhey, MD

## Slide 7: Mortality at 2 years by discharge mean gradient
Chart: 
<20 mmHg ≥20 mmHg | 223 221 130 130 | 201 111 | 163 89
No. subjects at risk
20.6%
10.1%
All-Cause Mortality %
Years Post Procedure
Log-rank P=0.008
David Holzhey, MD

## Slide 8: Multivariable predictors of 2-year mortality from time of discharge by failure mode
David Holzhey, MD
Variable | Hazard Ratio (95% CI) | P-value
Stenosis |  | 
Discharge MG ≥20 mmHg | 1.79 (0.87-3.66) | 0.11
STS Score, % | 1.06 (1.0-1.13) | 0.04
SAV tissue type |  | 
Bovine vs homograft | 0.37 (0.11-1.28) | 0.12
Porcine vs homograft | 0.86 (0.24-3.12) | 0.82
Regurgitation and combined |  | 
Discharge MG ≥20 mmHg | 2.64 (0.91-7.60) | 0.07
Age, years | 1.11 (1.03-1.19) | 0.005
Diabetes Mellitus | 2.21 (0.77-6.33) | 0.14
Prior Stroke or TIA | 4.41 (1.30-14.99) | 0.02

## Slide 9: Multivariable predictors of discharge mean gradient ≥ 20 mmHg
Variable | Odds Ratio [95% CI] | P-value
NYHA class III/IV | 0.62 (0.32-1.18) | 0.15
STS Score, % | 1.04 (0.99-1.09) | 0.09
TAV implant size |  | 
26 vs 23 mm | 0.20 (0.11-0.37) | 0.43
29/31 vs 23 mm | 0.08 (0.02-0.27) | <0.01
David Holzhey, MD

## Slide 10: Conclusions
This combined analysis shows good clinical and hemodynamic outcomes in these TAV in SAV patients.
Patients with a discharge mean gradient of ≥20 mmHg had a significantly higher rate of all-cause mortality at 2 years.
TAVI implant size of 23 mm and higher STS Score were predictors for discharge mean gradient ≥20 mmHg.
Discharge mean gradient ≥20 mmHg was significantly associated with 2-year mortality.
David Holzhey, MD