## Slide 1: The PARTNER 3 Bicuspid Registry for SAPIEN 3 TAVR in Low-risk Patients
Mathew R. Williams, MD & John G. Webb, MD
on behalf of the PARTNER 3 Trial Investigators

## Slide 2: Disclosures – Mathew Williams, MD
TVT: July 20-22, 2021
Within the past 36 months, I or my spouse/partner have had a financial interest/arrangement or affiliation with the organization(s) listed below:
Financial Relationship | Company
Research Support | Edwards, Medtronic, HLT, BSC
Other | Thubrikar, MID

## Slide 3: Background
Bicuspid aortic valve (BAV) morphology is a common congenital cardiac abnormality, affecting about 1% of Americans
Aortic stenosis (AS) is a frequent complication of bicuspid anatomy
Patients with bicuspid AS have historically been excluded from TAVR studies, however recent reports suggest that the use of SAPIEN 3 (S3) in high-risk patients is safe and effective
TAVR data for BAV patients at low surgical risk are limited

## Slide 4: Purpose
To prospectively study the 1-year safety and efficacy of the SAPIEN 3 TAVR system in selected patients with severe symptomatic AS and bicuspid anatomy who are at low surgical risk
To compare results to a clinically similar cohort of patients with tricuspid anatomy treated with SAPIEN 3 TAVR in PARTNER 3

## Slide 5: PARTNER 3 Study Design
1:1 Randomization
1000 Patients
Surgery
(bioprosthetic valve)
Symptomatic Severe AS
Follow-up: 30 days, 6 months, and annually through 10 years
PRIMARY ENDPOINT:
Composite of all-cause death, stroke, or CV rehospitalization at 1 year post-procedure
P3 TAVR Bicuspid Registry
P3 Bicuspid CAP TVT-Registry
Follow-up: 30 days and 1 year
All-cause death, stroke, valve-related rehospitalization,
and new PPMI
Low Risk/TF Assessment by Heart Team
(STS < 4%)
TAVR
(SAPIEN 3 THV)
Continued Access Protocol (CAP) for S3 TAVR
Managed by TVT registry
In addition to meeting all P3 Bicuspid eligibility criteria, patients had to be covered by Medicare
TAVR
(SAPIEN 3 THV)
TAVR
(SAPIEN 3 THV)

## Slide 6: Study Leadership
National Principal Investigators
Mathew R. Williams, MD, NYU Langone Medical Center, New York, NY
John G. Webb, MD, St. Paul’s Hospital, Vancouver, BC, Canada
Steering Committee
Howard Herrmann, Samir Kapadia, Susheel Kodali, Martin B. Leon, Michael J. Mack, Raj Makkar, Craig R. Smith (chair), Wilson Szeto, Vinod Thourani, John G. Webb
Data & Safety Monitoring Board
Cardiovascular Research Foundation, New York, NY; Joseph Carrozza, Jr., MD, chair
Clinical Events Committee
Cardiovascular Research Foundation, New York, NY; Ziad A. Ali, M.D., chair; Steven Marx, M.D., co-chair
CT Core Laboratory
The University of British Columbia; Jonathon Leipsic, MD, chair; Philipp Blanke, MD, chair
Echocardiographic Core Laboratory
Quebec Heart & Lung Institute (Laval University); Philippe Pibarot, DVM PhD, chair
Cardiovascular Research Foundation, New York, NY; Rebecca Hahn, MD, chair
Sponsor
Edwards Lifesciences, Irvine, CA

## Slide 7: Key Inclusion Criteria
Severe Calcific AS
AVA ≤ 1.0 cm2 or AVA index ≤ 0.6 cm2/m2
Jet velocity ≥ 4.0 m/s or mean gradient ≥ 40 mmHg, AND
  NYHA Functional Class ≥ 2, OR
  Abnormal exercise test with severe SOB, abnormal BP response, or arrhythmia, OR
  Asymptomatic with LVEF < 50%
Low Surgical Risk
STS < 4% or judgement of local heart team and national case review
board
Adjudicated by case review board
Bicuspid Aortic Valve Anatomy
Bicuspid valve as determined by the CT core lab

## Slide 8: Key Exclusion Criteria
Anatomic:
Aortic annulus diameter < 16 mm or > 28 mm (3D imaging)
Severe AR (> 3+) or MR (> 3+)
Severe LV dysfunction (LVEF < 30%)
Severe calcification of aortic valvar complex (esp. LVOT or raphe)
Vascular anatomy not suitable for safe femoral access
Complex CAD: ULM, Syntax score > 32, or not amenable for PCI
Low coronary takeoff (high risk for obstruction)
Clinical:
Acute MI within 1 month
Stroke or TIA within 90 days
Renal insufficiency (eGFR < 30 ml/min) and/or renal replacement Rx
Hemodynamic or respiratory instability
Frailty (objective assessment; > 2/4+ metrics)

## Slide 9: Study Methodology
Each patient reviewed (including imaging studies) by multidisciplinary heart team AND case review board
Baseline and 30-day neuro assessment in all patients; serial neurologist examinations and neuro-imaging for suspected neuro events
Baseline 3D cardiac imaging by CT or TEE
  Bicuspid anatomy determined by the CT core lab
100% adjudication of all major endpoint events (VARC-2 definitions when applicable)
  CEC for Bicuspid Registry and Tricuspid TAVR arm
  Independent medical reviewer for Bicuspid CAP

## Slide 10: Primary Endpoint
Composite endpoint of all-cause death, all stroke, and CV rehospitalization*
at 1 year post-TAVR
*Valve- or procedure-related and including HF

## Slide 11: Statistical Methods
Patients from both registries were pooled to form one BAV TAVR group
Propensity-matched analysis was used to compare the BAV TAVR group to the tricuspid TAVR arm of the PARTNER 3 trial
Descriptive statistics were presented as follows:
  Continuous variables summarized by mean ± SD
  Categorical variables summarized by count and percent
  Event variables summarized by Kaplan-Meier event rates and number of subjects with the event

## Slide 12: Covariates for Propensity Matching
Age
Renal disease
Sex
Atrial fibrillation
NYHA class
Prior cerebrovascular accident
Body mass index
Prior PCI
STS score
Aortic valve mean gradient
Diabetes mellitus
Effective orifice area
Peripheral vascular disease
Mitral insufficiency
Carotid disease
KCCQ score
Hypertension

## Slide 13: Patient Disposition
P3 Bicuspid Registry N=133
P3 Bicuspid CAP N=187
Not implanted (62):
Anatomical exclusion (43)
Enrollment closed (8)
Procedure not started (4)
Withdrew consent (3)
Not specified (2)
Medical exclusion (2)
Not implanted (89):
Anatomical exclusion (56)
Received commercial treatment (2)
Other (31)
N=71
N=98
P3 Tricuspid N=496
Total P3 Bicuspid N=169
148 Matched Pairs
Valve Implanted
Before Propensity Matching
After Propensity Matching

## Slide 14: Anatomical Exclusions
Reason(s) n (%) | Bicuspid Registry (N=43) | Bicuspid CAP (N=56)
Severe calcification of aortic valvar complex | 5 (11.6) | 31 (55.4)
Tricuspid aortic valve | 15 (34.9) | 0.0
Annulus size and/or ascending aorta diameter too large | 11 (25.6) | 14 (25.0)
AND severe calcification AND aortic stenosis not calcific AND risk of coronary obstruction | 3 (7.0) 1 (2.3) 1 (2.3) | 7 (12.5) 0.0 0.0
Risk of coronary obstruction | 1 (2.3) | 3 (5.4)
Other | 6 (14.0) | 1 (1.8)

## Slide 15: Baseline Patient Characteristics Matched
Variable (median or %) | Bicuspid (N=148) | Tricuspid (N=148) | Variable (median or %) | Bicuspid (N=148) | Tricuspid (N=148)
Age, years | 71.0 | 72.0 | Diabetes | 17.7% | 18.2%
Male | 58.1% | 60.1% | Mean gradient, mmHg | 49.0 | 50.1
BMI, kg/m2 | 28.0 | 27.6 | Aortic valve area, cm2 | 0.7 | 0.7
STS Score, % | 1.4 | 1.5 | MR ≥moderate | 3.5% | 2.7%
NYHA class III or IV | 29.7% | 26.4% | TR ≥moderate | 2.1% | 0.7%
KCCQ-OS | 72.7 | 72.9 | History of AFib | 4.7% | 1.4%
No significant differences among any variables.

## Slide 16: Baseline Morphology
Bicuspid Sievers Classification |  |  | 
 | Type 0 (No raphe) | Type 1 (1 raphe) | Type 2 (2 raphae)
Registry (N=71) CAP (N=98) | 14.1% 13.3% | 84.5% 86.7% | 1.4% 0%
% of patients
Figures adapted from Mylotte D et al. J Am Coll Cardiol. 2014 Dec, 64 (22) 2330-2339.

## Slide 17: Procedural Outcomes Matched
Outcome | Bicuspid (N=148) | Tricuspid (N=148) | p value
Conscious sedation | 68.9% | 62.2% | 0.27
Procedure duration (min) | 64.0 (46.0, 89.5) | 50.0 (36.0, 65.0) | <0.01
Fluoroscopy duration (min) | 13.3 (10.0, 19.8) | 12.0 (9.0, 16.0) | <0.01
Annulus area (mm2) | 486.0 (406.2, 541.6) | 457.1 (390.3, 516.3) | 0.01
Implanted valve size |  |  | 0.19
20 or 23 mm | 33.1% | 41.2% | 
26 or 29 mm | 66.9% | 58.8% | 
Oversizing (%) | 5.7±8.37 (N=147) | 8.0±8.82 (N=144) | 0.02
Procedure aborted | 0.0% | 0.0% | NA
Conversion to SAVR | 0.0% | 0.7% | >0.99
Procedural success | 100.0% | 99.3% | >0.99
mean ± SD, median (IQR), or %

## Slide 18: Primary Endpoint Matched
0
3
9
12
6
Months After Procedure
Death, Stroke, or Rehosp (%)
0
10
20
Tricuspid Bicuspid
Log Rank P = 0.80
10.9%
10.2%
Number at risk:
Tricuspid 148
Bicuspid 148
127
122

## Slide 19: Clinical Outcomes Matched
Outcome (KM estimate as %) |  | 30 Days |  |  | 1 Year | 
 | Bicuspid N=148 | Tricuspid N=148 | p value | Bicuspid N=148 | Tricuspid N=148 | p value
Death, Stroke, or Rehospitalization | 6.8% | 4.7% | 0.44 | 10.9% | 10.2% | 0.80
Death | 0.0% | 0.0% | NA | 0.7% | 1.4% | 0.58
Stroke | 1.4% | 1.4% | 0.99 | 2.1% | 2.0% | 0.99
Rehospitalization | 5.4% | 4.1% | 0.58 | 9.6% | 9.5% | 0.96
New pacemaker | 6.1% | 6.8% | 0.81 | 6.8% | 7.4% | 0.82

## Slide 20: Echocardiographic Findings Unmatched
0
0.5
1
1.5
2
0
20
40
60
Bicuspid CAP (site-reported) Bicuspid Registry (core lab-reported) Tricuspid (core lab-reported)
Mean Gradient (mmHg)
Baseline
30 Days
1 Year
Effective Orifice Area (cm2)
97|NA
71|65
484|459
95|NA
70|61
491|471
74|NA
68|65
470|447
No. of echos (MG|EOA):
Bicuspid CAP Bicuspid Registry Tricuspid

## Slide 21: Paravalvular Regurgitation
69.8
74.3
70.5
78.1
73.1
70.0
27.9
24.3
28.7
18.8
26.9
29.3
2.3
1.4
0.8
3.1
0.6
20
0
40
60
80
100
% of Patients
Bicuspid CAP N=70
Tricuspid Registry N=488
Unmatched
Bicuspid Registry N=86
30 Days
≥Moderate Mild None/Trace
Tricuspid Registry N=467
Bicuspid CAP N=67
Bicuspid Registry N=64
1 Year
Numbers reflect all available echos

## Slide 22: Paravalvular Regurgitation
70.5
72.4
64.6
77.6
69.1
68.6
26.9
25.9
35.4
20.7
30.9
30.7
2.6
0.7
20
0
40
60
80
100
% of Patients
Bicuspid CAP N=78
Tricuspid Registry N=147
Bicuspid Registry N=58
30 Days
≥Moderate Mild None/Trace
Tricuspid Registry N=140
Bicuspid CAP N=58
Bicuspid Registry N=55
1 Year
Matched
p = 0.119 p = >0.999
1.7 1.7
Numbers reflect all available echos

## Slide 23: NYHA Class III/IV Matched
0
1.4
0
10
20
30
40
P = 0.50
P = 0.60
29.7
26.4
P > 0.99
0.8 0.7
Tricuspid Bicuspid
% of Patients
Baseline
N=148 N=148
30 Days N=148 N=144
1 Year N=144 N=129

## Slide 24: KCCQ Overall Summary Score Matched
91.4
69.5
50
60
70
80
90
100
68.0
P = 0.54
88.4
P = 0.02
92.3
91.6
P = 0.50
Tricuspid Bicuspid
KCCQ OS Score
Baseline 30 Days
Bicuspid Tricuspid
N=147 N=148
N=148 N=145
1 Year
N=144 N=126

## Slide 25: Study Limitations
Results reflect only 1-year outcomes; longer-term data for P3
CAP are not available
Results apply only to the enrolled population of highly selected
bicuspid patients
Conclusions from this limited sample size may not be generalizable to all bicuspid morphologies
P3 CAP 30-day and 1-year echo data were site reported

## Slide 26: Conclusions
Patients with severe, symptomatic aortic stenosis, low surgical risk, and favorable bicuspid anatomy who underwent TAVR demonstrated:
Hemodynamic outcomes and adverse events that were comparable to a similar group of patients with tricuspid anatomy
Significant quality of life and symptomatic improvement
Longer-term follow up is necessary to confirm these results

## Slide 27: Thank You

## Slide 28: Back-up / Alternate Slides

## Slide 29: All-Cause Mortality Matched
0
3
6 9 12
Months After Procedure
All-Cause Mortality (%)
0
5
10
Tricuspid Bicuspid
Log Rank P = 0.58
1.4%
0.7%
Number at risk:
Tricuspid 148
Bicuspid 148
140
135

## Slide 30: All Stroke Matched
0
3
9
12
6
Months After Procedure
Tricuspid Bicuspid
All Stroke (%)
0
10
2.1%
2.0%
Log Rank P = 0.99
Number at risk:
Tricuspid 148
Bicuspid 148
139
133
5

## Slide 31: Rehospitalization Matched
0
3
9
12
6
Months After Procedure
Rehospitalization (%)
0
10
20
Tricuspid Bicuspid
Log Rank P = 0.96
9.6%
9.5%
Number at risk:
Tricuspid 148
Bicuspid 148
127
123

## Slide 32: New Permanent Pacemaker Matched
0
3
9
12
6
Months After Procedure
New PPM (%)
0
10
20
Tricuspid Bicuspid
Log Rank P = 0.82
7.4%
6.8%
Number at risk:
Tricuspid 148
Bicuspid 148
130
126