## Slide 1: Incidence, Mechanisms, Predictors, and Evolution of Severe Peri-Device Leaks: New Insights from Amulet IDE Trial
Author(s): Dhanunjaya Lakkireddy Jens Erik Nielsen-Kudsk, Stephan Windecker, David Thaler, Matthew J. Price, Alok Gambhir, Nigel Gupta, Konstantinos Koulogiannis, Leo Marcoff, Anuj Mediratta, Christopher R. Ellis

## Slide 2: What we know about PDL
Various shapes and sizes of the LAA may result in peri-device leak (PDL) which have been associated with worse clinical outcomes
Amulet maintains significantly fewer severe PDLs (>5mm) than Watchman through 12 months
Amulet provides superior complete closure

## Slide 3: Objective
To present the incidence and predictors of severe PDL (>5mm) in the Amulet IDE trial
Identify the mechanistic factors of severe PDL and evolution through 18 months

## Slide 4: Methods
PDL Definition
PDL was graded as the single largest jet visualized around the device on TEE from a minimum of 3 Doppler view by an independent core lab
Severe PDL was defined as (>5mm) based on Saw et al1 and Lakkireddy et al2
Analysis Outcomes
Incidence of severe PDL at 45 days and 12 months
Implant mechanism resulting in severe PDL
Anatomical, procedural, and device-based predictors of severe PDL using propensity score matching analysis
Evolution and temporal observations of severe PDL through 18 months
1) Saw et al. 2017. JACC Cardiovasc Interv. 10(4):391-399.
2) Lakkireddy et al. Circulation 2021;144:1543-1552.

## Slide 5: Significantly More Severe PDLs with Watchman Device Through 12 Months
Analysis includes only the first-time identification of a severe PDL through 12-month TEE follow-up p<0.01 using paired t-test
Chart: Incidence of Severe PDLs Through 12 Months
| Amulet | Watchman
PDL Through 12 Months | 10 | 32
N=903
N=885

## Slide 6: No Anatomical Limitations with Amulet Occluder
Larger LAA anatomical features predicted severe PDL in the Watchman device with no anatomical predictors in the Amulet occluder
Anatomical Variable | Severe PDL (n=10) | No PDL (n=46)
Orifice diameter, mm | 21.0 ± 6.4 (9) | 18.6 ± 4.7 (37)
Landing zone diameter, mm | 17.7 ± 4.7 (9) | 15.6 ± 3.7 (37)
Depth, mm | 27.8 ± 7.7 (9) | 24.9 ± 6.0 (37)
Ovality index, ratio | 1.5 ± 0.5 (9) | 1.4 ± 0.3 (36)
Chicken wing (vs. others) | 3 (9) | 8 (38)
Multi-lobed LAA | 3 (9) | 11 (35)
*Indicates p<0.05
Total evaluable TEEs for each variable is listed in parentheses
Amulet
Anatomical Variable | Severe PDL (n=32) | No PDL (n=38)
*Orifice diameter, mm | 20.5 ± 5.1 (26) | 18.1 ± 3.1 (32)
*Landing zone diameter, mm | 16.8 ± 3.4 (26) | 14.9 ± 2.9 (32)
*Depth, mm | 26.9 ± 4.6 (28) | 23.4 ± 4.8 (33)
Ovality index, ratio | 1.3 ± 0.3 (26) | 1.4 ± 0.2 (32)
Chicken wing (vs. others) | 6 (29) | 10 (31)
Multi-lobed LAA | 12 (27) | 15 (29)
Watchman
Notes: 6

## Slide 7: Implant Experience Related Predictors with Amulet Occluder
Increased procedural time and number of devices used predicted severe PDL in the Amulet occluder with no procedural predictors in the Watchman device
Procedural Variable | #Severe PDL (n=10) | No PDL (n=46)
*Procedure duration, min | 55.5 ± 26.4 (10) | 36.3 ± 19.9 (46)
Contrast volume, cc | 117.2 ± 102.7 (10) | 80.2 ± 50.5 (46)
Number of recaptures | 2.2 ± 2.4 (10) | 1.3 ± 2.0 (46)
*Number of devices | 1.7 ± 0.9 (10) | 1.2 ± 0.4 (46)
Smoke in LAA | 4 (9) | 14 (41)
LAA flow velocity, cm/s | 16.6 ± 14.8 (6) | 33.6 ± 28.2 (23)
Afib rhythm (vs. sinus) | 7 (10) | 15 (36)
Amulet
Watchman
Procedural Variable | Severe PDL (n=32) | No PDL (n=38)
Procedure duration, min | 29.4 ± 16.1 (32) | 27.3 ± 19.7 (38)
Contrast volume, cc | 89.2 ± 55.3 (31) | 67.4 ± 51.4 (37)
Number of recaptures | 0.7 ± 1.0 (32) | 0.6 ± 1.3 (38)
Number of devices | 1.3 ± 0.5 (32) | 1.3 ± 0.5 (38)
Smoke in LAA | 8 (30) | 14 (35)
LAA flow velocity, cm/s | 21.6 ± 30.2 (15) | 16.9 ± 26.5 (22)
Afib rhythm (vs. sinus) | 10 (22) | 12 (33)
*Indicates p<0.05
#All 10 severe PDLs were from implanters who had less than 10 Amulet cases
Total evaluable TEEs for each variable is listed in parentheses
Procedure times cannot be directly compared between device groups due to additional angiogram views and measurements in the Amulet group required by the protocol
Notes: 7

## Slide 8: Implant Procedure Related Predictors with Both Devices
Improper device placement was a severe PDL predictor in both groups with incorrect device size and lobe off-axis also observed predictors of Amulet occluder
Device Variable | Severe PDL (n=10) | No PDL (n=46)
Depth of implant | 3.0 ± 3.7 mm (9) | 5.1 ± 5.5 mm (45)
Distance between disc and lobe | 5.5 ± 3.7 mm (10) | 3.6 ± 2.0 mm (45)
Max compression of device | 18.1% (9) | 18.5% (38)
*Incorrect device size | 8 (9) | 11 (36)
Undersized | 7 | 9
Oversized | 1 | 2
*Lobe off-axis | 5 (7) | 2 (20)
*Improper device placement | 4 (10) | 4 (46)
Too distal | 1 | 2
Too proximal | 3 | 2
Amulet
*Indicates p<0.05
#A majority of Watchman devices in this analysis were incorrectly sized although there was no significant differences based off PDL status
Total evaluable TEEs for each variable is listed in parentheses
Device Variable | Severe PDL (n=32) | No PDL (n=38)
Depth of implant | 10.1 ± 7.5 mm (30) | 8.4 ± 6.9 mm (36)
Max compression of device | 25.6% (27) | 22.6% (35)
#Incorrect device size | 16 (26) | 19 (32)
Undersized | 6 | 4
Oversized | 10 | 15
*Improper device placement | 24 (30) | 2 (38)
Too distal | 6 | 0
Too proximal | 1 | 1
Off-axis | 17 | 1
Watchman

## Slide 9: Greater Number of Adverse Events and Unresolved PDLs with the Watchman Device
Watchman device had a greater number of adverse events, deaths, and unresolved PDLs by 18 months compared to the Amulet occluder
 | Amulet (n=10) | Watchman (n=32)
Medication at time of PDL identification |  | 
Oral anticoagulation | 4 | 25
Antiplatelet | 6 | 7
Adverse events after PDL | 2 | 6
Major bleeding | 1 | 4
Ischemic stroke | 1 | 1
Transient ischemic attack | 0 | 1
Deaths after PDL | 0 | 4
*Resolution status |  | 
Unresolved (>5mm) | 3 | 11
Partially resolved (≤5mm) | 4 | 5
Fully resolved (0mm) | 1 | 3
*Total evaluable TEEs for Amulet (n=8) and Watchman (n=19) for resolution status
Severe PDL Evolution Through 18 Months

## Slide 10: New Insights from Severe PDL Predictor Analysis
The dual-seal Amplatzer Amulet LAA occluder provides clinically relevant closure with newly discovered PDL predictors in the Amulet IDE trial
LAA anatomy is a predictor of PDL with Watchman and perhaps the device does not treat all anatomies well, specifically larger LAA dimensions
No LAA anatomical limitations with Amulet with complete closure in complex anatomies
Expected learning curve with the Amulet procedure should further improve closure results
Majority of Watchman device severe PDLs did not resolve over time with increased number of adverse events and deaths through 18 months